© Copyright The University of Montana, Rural Institute, MonTECH

Cleaning and Disinfection Procedures

# MonTECH’s Cleaning and Disinfection Procedures

Contents

COVID-19 Review of Cleaning Procedures	2

General Cleaning and Disinfection Instructions	4

Cleaning Station Surfaces and Office: Cleaning & Disinfection Procedures	8

Cleaning Office and High Touch Surfaces: Cleaning & Disinfection Procedures	9

Assistive Technology Not Otherwise Covered by Specific Instructions: Cleaning & Disinfection Procedures	10

Electronics - Tablets, Computers, Lens Magnifiers, Digital Magnifiers and Smart Devices: Cleaning & Disinfection Procedures	12

Computer Keyboards, Mice and Stands: Cleaning & Disinfection Procedures	14

Communication Devices: Cleaning & Disinfection Procedures	16

Switches and Switch Accessories: Cleaning & Disinfection Procedures	18

Equipment for Eating and Food Preparation: Cleaning & Disinfection Procedures	20

Telepresence Robots: Cleaning & Disinfection Procedures	22

Headsets, Microphones, and FM Systems: Cleaning & Disinfection Procedures	24

Rifton Pacer Gait Trainers, Creepster Crawler: Cleaning & Disinfection Procedures	26

Crocodile, Pony, and Mustang Gait Trainers: Cleaning & Disinfection Procedures	28

Rifton TRAM: Cleaning & Disinfection Procedures	30

Tumble Forms 2 Products, Special Tomato Soft Touch Floor Sitter with Material or Wooden Wedge: Cleaning & Disinfection Procedures	32

Special Tomato Height Right, Keekaroo Height Right and Stokke Tripp Trapp Chairs (wooden with straps): Cleaning & Disinfection Procedures	34

Kaye Tilting Benches, Classroom Positioner Chair: Cleaning & Disinfection Procedures	36

Wenzelite Chair, Prime Engineering Stander and Mulholland Standers:	38

Rifton Supine and Prone Stander: Cleaning & Disinfection Procedures	40

EasyStand Standing Frames: Cleaning & Disinfection Procedures	42

Rifton Large Dynamic Stander and TumbleForms Tristander 45:	44

Rifton Pacer Accessories: Cleaning & Disinfection Procedures	46

Bumbo Baby Seat and Rounded Sit-Me-Up Seat: Cleaning & Disinfection Procedures	47

Birillo, LadyBug: Cleaning & Disinfection Procedures	48

Firefly GoTo Seat: Cleaning & Disinfection Procedures	50

Manual Wheelchairs, Walkers, Rollators, Knee Scooters, Canes and Crutches: Cleaning & Disinfecting Procedures	51

Adaptive Strollers: Cleaning & Disinfecting Procedures	52

AmTryke Early Intervention Tryke: Cleaning & Disinfection Procedures	54

# COVID-19 Review of Cleaning Procedures

## Purpose

Review standard cleaning and sanitization procedures in light of COVID-19 pandemic and indications that the virus may live from a few hours to a few days on a variety of surfaces. Note this research is emerging and the exact length of time the virus remains may not be known.

“SARS-CoV-2 remained active on plastic and stainless-steel surfaces for two to three days under the conditions in this experiment. It remained infectious for up to 24 hours on cardboard and four hours on copper. The virus was detectable in aerosols for up to three hours. These times will vary under real-world conditions, depending on factors including temperature, humidity, ventilation, and the amount of virus deposited.” NIH March 24, 2020 Study suggests new coronavirus may remain on surfaces for days

## Sources Used to Revise Cleaning Procedures

- Manufacturer’s recommendations: Goal is to ensure compatibility of cleaning and disinfection chemicals with materials on equipment. Preventing damage to our equipment is important not only to protect our financial investment in equipment, but also to ensure safe use of equipment (e.g. straps or harnesses on adaptive equipment). Not all equipment will have cleaning recommendations or may have cleaning recommendations that do not kill infectious agents. In that event, we will use the gentlest cleaning procedures that meet cleaning and disinfection guidelines.

- AT3 Webinar: A Framework for Cleaning & Sanitization: Back to Basics & Quick Tips 11/6/19

- Guideline for Disinfection and Sterilization in Healthcare Facilities (CDC, 2008; updated 2018): *This is the guidelines AT3 webinar recommends we follow at all times (i.e. pandemic or no pandemic)

  - Perform low-level disinfection for noncritical patient-care surfaces (e.g., bedrails, over-the-bed table) and equipment (e.g., blood pressure cuff) that touch intact skin

  - Process noncritical patient-care devices using a disinfectant and the concentration of germicide listed in Table 1 (Ethyl or isopropyl alcohol (70-90%) for low-level disinfection of hard, smooth surfaces).

  - Disinfect noncritical medical devices (e.g., blood pressure cuff) with an EPA-registered hospital disinfectant using the label’s safety precautions and use directions.

  - 3/31/20: Marlena, Dave, and Claire feel that using a hospital grade disinfectant whenever it will not damage the equipment is the safest alternative. For delicate technology that is not compatible with a disinfectant wipe, use a 70% isopropyl alcohol wipe.

- CDC Guidelines Environmental Cleaning and Disinfection Recommendations from Interim Recommendations for US Community Facilities with Suspected/Confirmed Coronavirus Disease 2019

  - If surfaces are dirty, they should be cleaned using a detergent or soap and water prior to disinfection.

  - For hard, non-porous surface disinfection:

    - diluted household bleach solutions and most common EPA-registered household disinfectants should be effective.

    - Products with EPA-approved emerging viral pathogens claims are expected to be effective against COVID-19 based on data for harder to kill viruses. Follow the manufacturer’s instructions for all cleaning and disinfection products (e.g., concentration, application method and contact time, etc.).

  - For soft (porous) surfaces such as carpeted floor, rugs, and drapes, remove visible contamination if present and clean with appropriate cleaners indicated for use on these surfaces. After cleaning:

    - If the items can be laundered, launder items in accordance with the manufacturer’s instructions using the warmest appropriate water setting for the items and then dry items completely.

    - Otherwise, use products with the EPA-approved emerging viral pathogens claims (examples at this link pdf) that are suitable for porous surfaces

  - For Electronics

    - Follow the manufacturer’s instructions for all cleaning and disinfection products.

    - Consider use of wipeable covers for electronics.

    - If no manufacturer guidance is available, consider the use of alcohol-based wipes or sprays containing at least 70% alcohol to disinfect touch screen

  - EPA List N: Disinfectants to Use Against SARS COV 2 List N includes products that meet EPA’s criteria for use against SARS-CoV-2, the novel coronavirus that causes the disease COVID-19

- OSHA Standards for bloodborne pathogens (BBP, 29CFR 1910.1030): Treat all used equipment like it carries infectious agents.

  - Standard precautions (SP): hand hygiene and PPE for all contaminated equipment.

  - SP Applies to all patients even if not suspected or known to be infections.

- Shipping and Receiving:

  - World Health Organization (WHO) states that the likelihood of infected person contaminating packages is low and states it is safe to receive a package from an area where COVID has been reported.

  - United States Postal Service Media Statement 4/2/20: “Importantly, the CDC (https://www.cdc.gov/coronavirus/2019-ncov/faq.html), the World Health Organization (https://www.who.int/news-room/q-a-detail/q-a-coronaviruses), as well as the Surgeon General have indicated that there is currently no evidence that COVID-19 is being spread through the mail.”

  - CDC’s Guidelines for Mail and Parcel Drivers

# General Cleaning and Disinfection Instructions

## When & Where to Clean

All equipment will be cleaned and disinfected after each use before returning to storage shelves.

### Equipment Loaned to Consumers

- All returned loan items must be cleaned and disinfected in the designated cleaning station before re-shelving.

  - If the Billings office receives equipment that is to be shipped to the Missoula office, it will be cleaned at the Billings office and package will contain a note stating that cleaning and disinfection is complete.

- Surfaces in the cleaning station (tables, carts, tubs, etc.) and tools (vacuum head, tweezers, pick, etc.) must be cleaned and disinfected after each “batch” of equipment is cleaned (e.g. all items returned from the same user).

- All outgoing loan items must be inspected prior to distributing to consumers.

  - If the item is visibly dirty it must be cleaned and disinfected prior to distribution.

  - If the item is visibly clean it must be disinfected prior to distribution.

### Equipment Used for Demonstrations or Office Use

- All equipment used for demo, testing or in-office use must be inspected prior to consumer contact.

  - If the item is visibly dirty it must be cleaned and disinfected prior to consumer contact.

  - If the item is visibly clean it must be disinfected prior to consumer contact.

- All equipment used for demo, testing or in-office use must be cleaned and disinfected after each use before re-shelving. Demo equipment may be cleaned and disinfected in place or transported to the designated cleaning station.

- Surfaces used in the demo (tables, desks, shelves) should be cleaned and disinfected after each appointment.

- A hand washing station and hand sanitizer will be available to staff and consumers during demonstrations.

### Equipment Used for Trainings and Public Awareness Events

- When we can safely gather in crowds without social distancing restrictions, equipment can be trialed by multiple people during an event.  Until that time, equipment will only be handled by staff in “show and tell” format.

- All equipment used at these events will be cleaned and sanitized at the end of the event. This can be completed in place or brought back to the cleaning station in the office before being re-shelved.

- Hand sanitizer will be available to staff and consumers during events.

## Cleaning Station

### Missoula

Dirty items will be placed on “equipment return” cart in entry for transport to the storage room (former conference room) or to designated cleaning room (former restroom). Large items will be placed in the conference room immediately upon return from consumer.

If water is needed for cleaning, fill a plastic tote or tub with warm soapy water and fill a second plastic tote or tub with clean “rinse” water in cleaning area. Do not clean or disinfect items at the kitchen/break room sink. If the handwashing items too large to fit in the portable washing machine use the utility sink in the cleaning room. The utility sink must be disinfected before and after use.

All surfaces used to hold “dirty” items will be cleaned and disinfected after each use including plastic totes, carts, desk surface, etc.  After cleaning and disinfection, items should only be placed in a “clean” plastic tote or cart to re-shelve and should not be placed back in a tote used to transport dirty items until the tote has been cleaned and disinfected.

If items are returned that cannot be properly cleaned and disinfected at this time due to lack COVID-19 guidance for certain materials or lack of disinfection supplies, items will be tagged “Need Cleaning and Disinfection” and stored in the conference room.

### Billings

Dirty items will be placed on rolling carts located near office bookshelf. Items that do not fit on the cart (e.g. scooter, TRAM, etc.) will be left in the center of office. Clean and disinfect all items on the section of desk marked with tape. If items require water (e.g. food prep equipment), fill cleaned and disinfected bucket or basin in bathroom or sink and use rolling cart to transport to office cleaning area.

All surfaces used to hold “dirty” items will be cleaned and disinfected after each use including carts, desk surface, buckets, basins, etc.  After cleaning and disinfection, items should only be placed on a “clean” cart to re-shelve and should not be placed back on a cart used to transport dirty items until the cart has been cleaned and disinfected.

If items are returned that cannot be properly cleaned and disinfected at this time due to lack COVID-19 guidance for certain materials or lack of disinfection supplies, items will be tagged “Need Cleaning and Disinfection” and stored in basement storage area.

## General Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on disposable nitrile gloves and personal face mask. Gloves protect your health by preventing contact with germs that may be on equipment and from the chemicals in the cleaning and disinfection products.

- Prior to cleaning and disinfection, items with hinges should be opened (e.g. open stand on portable video magnifier) and equipment should be disassembled as much as possible where designed to be disassembled (remove cases, covers, etc.).

- Clean prior to disinfecting following item-specific directions. Cleaning removes dirt and grime which can contain infectious bacteria, virus, yeast, etc. If devices are not cleaned first, the disinfection process will not be effective.

- Disinfect items following item-specific directions. Disinfecting the equipment kills infectious bacteria, virus, yeast, etc. Disinfectant wipes contain a chemical that requires contact time with infectious agents to work. Follow the longer time listed either: on chemical instructions label or List N (this will be listed in MonTECH’s cleaning procedures). If the item dries prior to the required contact time, use new, fresh wipes to keep the surface wet for the required contact time. Some products allow use of the same type of wipe for both cleaning and disinfection. A new, fresh wipe must always be used for disinfection. Contact time is not important during cleaning, but is for disinfection. Cleaning and disinfection rules apply to all accessories, cases and packaging handled by consumer.

- Put cleaned item on “clean” cart to await re-shelving (see last step).

- Remove gloves following doffing procedures that prevent contamination of your hands and throw in trash. Gloves may be re-used while cleaning the same “batch” of items, such as multiple loan items returned from the same user, or multiple items used in a demonstration or training, etc. New gloves should be used before cleaning the next “batch” of equipment or after handling any item that is heavily soiled. Changing gloves prevents cross contamination between items.

- Wash your hands with soap and water after cleaning and disinfection. This further protects you from contacting infectious agents.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## Training

All staff will be trained annually on the cleaning and disinfection procedures. New staff will be trained during orientation. If the procedures are updated as new information becomes available, all staff will be trained on updates.

## Storage

### Material Safety Data Sheets (MSDS)

An MSDS for every cleaning and disinfecting chemical used will be posted near chemical storage. MSDS contain safety information, including necessary personal protective equipment (PPE) and first aid procedures, if necessary.

### Location

All cleaning and disinfecting supplies will be stored in a locked cabinet in the designated cleaning area.

### Bottles and Labels

All chemicals will be kept in original bottles with manufacturer labels. Any additional bottles (e.g. water spray bottle) will be labeled with permanent marker.

## Cleaning Products and Applied Use

Substitution of cleaning products may be necessary due to delivery delays or supply shortages. If cleaning product listed in specific equipment procedures is unavailable and a substitute is used always refer to label instructions for contact time required to achieve disinfection. Contact time will vary for each cleaning product. Medical grade products are preferred for disinfection whenever possible.

Products applied to any item that may be placed in the mouth must be rinsed with the exception of the Purell Foodservice Surface Sanitizer. Items of this nature that cannot be rinsed, such as plush toys, may not be distributed to consumers.

### General cleaning (not disinfecting)

Dish Soap

Lysol Brand Clean & Fresh Multi-Surface Cleaner (diluted)

### Hard Surface Disinfecting

PDI Sani-Cloth AF3 Germicidal Disinfecting Wipes (medical grade)

Purell Healthcare Surface Disinfectant (medical grade)

Clorox Healthcare Bleach Germicidal Cleaner (medical grade)

Lysol Brand Clean & Fresh Multi-Surface Cleaner

Clorox Bleach-Free Disinfecting Wipes

70% Isopropyl Alcohol or Alcohol Prep Pads

### Lens, Tablet and Digital Screen Device Disinfecting

PDI Easy Screen Cleaning Wipes

70% Isopropyl Alcohol or Alcohol Prep Pads

Clorox bleach-free disinfecting wipes (ONLY for iPads and Apple watches)

### Food Contact Surface Disinfecting (including utensils)

Purell Foodservice Surface Sanitizer - NOT Purell Healthcare Surface Disinfectant (bottles look similar).

Bleach (diluted)

### Soft Surface Sanitizing

Purell Healthcare Surface Disinfectant (medical grade)

Purell Foodservice Surface Sanitizer

Lysol Disinfecting Spray

Phenomenal Citrus Hospital Disinfectant

# Cleaning Station Surfaces and Office: Cleaning & Disinfection Procedures

## Examples of Equipment

Counters

Desks

Shelves

Bookcases

Carts

Buckets

Dish washing basin

Drying Rack

## Supplies Needed

- Hard surface disinfectant

- Shop towels or sponges

- PPE: gloves and mask

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Clean prior to disinfecting:

  - Use a spray cleaner and shop towel (or disinfectant wipes) to remove any chunks or visible dirt and grime on the entire surface (table, desk, etc.). Throw towels or wipes used for cleaning into the trash.

- Disinfect:

  - Apply hard surface disinfectant spray or a new disinfectant wipe to the entire surface. Reapply spray or wipe as needed to achieve the required contact time. If the surface is used for food preparation a clean water rinse is required after disinfection.

- After cleaning or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

# Cleaning Office and High Touch Surfaces: Cleaning & Disinfection Procedures

This procedure addresses routine cleaning of high-touch surfaces in the office.  If there is a client or staff with COVID-19 in the office, follow CDC guidelines for cleaning and disinfecting the office when someone is sick.

## When

- At the end of each day or when visibly dirty.

## Who

- Personal work station: Everyone cleans their own

- Common Areas:

  - Missoula: Maddy Musson or assigned

  - Billings: Marlena Lanini or assigned

- This procedure assumes toilets cleaned daily by University cleaning staff.

## What

- Door handles, light switches, desk, break area counter and faucet

- High touch office equipment at your desk such as keyboards, mouse, office phone

- Electronic devices used frequently during day such as tablet, laptop, touch screen, cell phone, printer/copier

## Equipment Types and Appropriate Cleaning Procedures

- Door handles, light switches, desk, break area counter and faucet: Follow Cleaning Station Surfaces and Office: Cleaning & Disinfection Procedures

- Equipment at your desk (Keyboards, mouse, office phone): Follow Hard Surface Computer Accessories and Smart Devices: Cleaning & Disinfection Procedures

- Tablet, laptop, touch screen, cell phone, printer/copier: Follow Tablets, Laptops, and Video Magnifiers: Cleaning and Disinfection Procedures

## Personal Protective Equipment Needed

- If you are cleaning only your own workspace you may determine your own PPE.

- If you are cleaning common office areas or surfaces touch by multiple people gloves should be worn.

# Assistive Technology Not Otherwise Covered by Specific Instructions: Cleaning & Disinfection Procedures

## Examples of Equipment

Anything made of hard, non-porous materials like hard plastic, glass, metal, etc. such as mounting systems, reachers, adaptive scissors, handwriting aids, telephones, alert systems, etc. Items can also include soft, porous surfaces such as Velcro, wrist straps, covers, cases, dense foam handle bars, grips, etc.

This does NOT apply to items with:

- Leather, untreated wood, soft open cell foam that doesn’t have a covering (CANNOT be loaned at this time)

- Food contact surfaces, magnifiers, or anything electronic; please see item specific instructions

- Adaptive Equipment; please see item specific instructions

## Supplies Needed

- Hard surface disinfectant

- Soft surface disinfectant

- Water basin, dish soap and disposable sponges

- Cotton swabs, shop towels

- PPE: gloves and mask

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Unplug all external power sources. Remove batteries for devices with removable batteries.

- Open any hinged parts and disassemble anything designed to be taken apart (e.g. modular hose pieces).

- Clean prior to disinfecting:

  - Remove any visible dirt and grime from hard surfaces on the entire device and packaging using a shop towel dipped in soapy water or a disposable disinfectant wipe.

  - Spot clean any soft surfaces that can’t be removed using a sponge or shop towel with dish soap and water. Rinse soap with clean water wipe. If item can be machine washed or submerged in a basin of water and soap, do so. Let air dry completely before disinfection step.

  - If there are tight areas that cannot be reached with wipe, wrap a wipe over the tip of a cotton swab.

  - Throw wipes, sponges and shop towels used for cleaning into the trash.

- Disinfect:

  - Apply hard surface disinfectant to the entire device including the screen, covers, accessories, packaging and all cords. Reapply as needed to achieve the required contact time.

  - Spray soft surfaces and cases with soft surface disinfectant. Reapply as needed to achieve the required contact time.

- Put cleaned and disinfected item on “clean” cart to await re-shelving (see last step).

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

How to Clean your AbleNet Device

# Electronics - Tablets, Computers, Lens Magnifiers, Digital Magnifiers and Smart Devices: Cleaning & Disinfection Procedures

## Examples of Equipment

iPad & Android Tablets

Laptop Computers including Microsoft Surface

Smart Watches (Apple Watch, Octopus, Time Timer)

Smart Home: Google Home, Amazon Echo (older hard plastic versions), plugs, hubs, etc.

Notetaking/Smart Pens**

C-Pen Reader

Webcams

UbiDuo

Eschenbach MenasLux

Desktop Magnifier

## Supplies Needed

- Disposable, scratch-free, soft, lint-free wipes such as KimTech Delicate Task Wipers

- Touch screen disinfectant: PDI Easy Screen Cleaning Wipes (70% isopropyl Alcohol)

  - Appropriate alternative:

    - 70% isopropyl alcohol prep wipes after gently squeezing excess liquid out into trash.

    - Liquid 70% isopropyl alcohol applied to KimTech Delicate Task Wiper

    - Clorox bleach-free disinfecting wipes (ONLY for iPads and Apple watches)

- Soft surface disinfectant

- Basin, dish soap and shop towels or sponges

- Electronics/Keyboard vacuum

- Blu-Tack adhesive putty, electronics cleaning gel or similar product

- PPE: gloves and mask

## Warnings

- Use only a soft, lint-free cloth. Avoid abrasive cloths, towels, paper towels, or similar items

- Avoid excessive wiping, which might cause damage.

- Don't get moisture into any openings. Wipes should be slightly damp, not wet or dripping.

- Don't use aerosol sprays, bleaches, or abrasives.

- Don't spray any cleaners directly onto the item. Always apply cleaner to cloth.

- Do not use bleach, ammonia or abrasive disinfectants. Check product compatibility if considering any new disinfectants.

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Unplug all external power sources.

- Open up any hinged parts (e.g. stand on video magnifiers) and disassemble anything designed to be taken apart (e.g. iPad cases).

- Clean prior to disinfecting:

  - If visible dirt and chunks are present in the laptop keyboard or other crevices, use keyboard vacuum to remove particles. If speakers contain visible grime, use a small amount of Blu-Tack adhesive putty or electronics cleaning gel to pull out grime. If there are tight areas that cannot be reached with wipe, use tip of a cotton swab to scrub these areas. Throw used putty and gel in trash. Clean equipment packaging (boxes, trays, etc.) as needed to remove dirt or debris.

  - Gently wipe screens and surfaces with a PDI Easy Screen Cleaning wipe (or alternative) to remove any visible dirt and grime.

  - Spot clean cases, covers with soft backing and soft surface packaging using a shop towel with soapy water followed by wipe with just water to remove soap. Allow to completely air dry before disinfection.

- Disinfect:

  - Gently wipe the entire device including the screen and all cords with either Clorox bleach-free disinfecting wipes (Apple products ONLY), PDI Easy Screen Cleaning wipes or an acceptable alternative. This includes the Alcantara (leather-like fabric) on the Microsoft Surface keyboard.  Ensure the wipe remains damp while wiping the entire device; more than 1 wipe may be needed if it dries out.

  - Soft Cases, covers with soft backing and item packaging: After the item has dried completely, spray with soft surface disinfectant and reapply as needed to achieve the required contact time.

- Place cleaned and disinfected items on “clean” cart to await re-shelving.

- Sanitize keyboard vacuum if used. Spray electronics vacuum brush with soft surface disinfectant and reapply as needed to achieve required contact time.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Use clean cart to transport items to storage and re-shelve.

## References

How to Clean your Apple Products

How to Clean your Microsoft Surface

How to Clean your Lenovo

Freedom Scientific: How to Clean your Topaz and Ruby

Enhanced Vision: How to Clean your Pebble and Amigo

- Enhanced vision does not recommend cleaning with anything except water. However, this will not disinfect the device, so the gentlest cleaning and disinfection procedures are used.

Smart Speakers: Both Google Home and Alexa recommend cleaning with dry cloth. However, this will not clean, nor disinfect the device, so the gentlest cleaning and disinfection procedures are used.

**(Livescribe controls must be laminated to be disinfected; send only sheets of paper, not entire notebook)

# Computer Keyboards, Mice and Stands: Cleaning & Disinfection Procedures

## Examples of Equipment

Keyboards

Mice

Stands

Low-tech pointing devices (head pointers, mouth sticks)

## Supplies Needed

- Blu-Tack adhesive putty, electronics cleaning gel or similar product

- Hard surface disinfectant

- Cotton swabs

- Electronics/Keyboard vacuum

- Soft Surface disinfectant

- Basin, dish soap and shop towels or sponges

- PPE: gloves and mask

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Unplug all external power sources. Remove batteries.

- Open up any hinged parts (e.g. stands) and disassemble anything designed to be taken apart (e.g. Ball on Kensington Trackball).

- Clean prior to disinfecting:

  - Use the keyboard/electronics vacuum to remove any loose particles present in keyboards or other crevices (e.g. mouse scroll wheel, etc.). Use a disposable wipe, electronics cleaning gel or adhesive putty to remove any chunks or any visible dirt and grime on the entire device including hard surface packaging. A cotton swab may be used to reach small spaces.

  - Throw wipes, gel, putty or cotton swabs used for cleaning into the trash.

  - For mice that have parts that go in user’s mouth (e.g. Glassouse, QuadJoy, etc.), the disposable mouth piece should be removed and thrown away. A new one will be sent for the next loan.

  - Spot clean soft cases with soapy water and a shop towel or sponge followed by wipe with just water to remove soap. Allow to completely air dry before disinfection.

- Disinfect device:

  - Using a new hard surface disinfectant wipe, wipe the entire device including any accessories, cords and hard surface packaging. Wrap a wipe around a cotton swab to reach small spaces. Reapply a new disinfectant wipe as needed to achieve required contact time.

  - Spray soft cases with soft surface disinfectant. Reapply as needed to achieve the required contact time.

- Mouth sticks and any other device that does not have a disposable mouth piece, must be rinsed with water after disinfection to remove any residue.

- Put cleaned and disinfected item on “clean” cart to await re-shelving (see last step).

- Disinfect keyboard vacuum, if used, with soft surface disinfectant and reapply as needed to achieve required contact time.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

# Communication Devices: Cleaning & Disinfection Procedures

## Examples of Equipment

Tobii I-12, I-12+, Indi

PRC Accent, PRiO

Big Mac, Talking Brix, Sequencer

Voice amplifiers (ChatterVox, Spokeman Luminaud)

AbleNet QuickTalker Freestyle

QuickTalker and GoTalk*

## Supplies Needed

- Blu-Tack adhesive putty, electronics cleaning gel or similar product

- Hard surface disinfectant

- Screen cleaner

- Soft Surface disinfectant

- Basin, dish soap and shop towels or sponges

- Cotton swabs

- PPE: gloves and mask

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Unplug all external power sources. Remove batteries for devices with removable batteries.

- Open up any hinged parts (e.g. stand on AAC) and disassemble anything designed to be taken apart.

- Clean prior to disinfecting:

  - Use a disposable wipe, slightly damp shop towel or sponge to remove any chunks, and any visible dirt and grime on the entire device (including the screen) and all accessories (e.g. cases, keyguards, cords, laminated paper inserts from GoTalks and QuickTalkers).

  - If speakers are visibly dirty, use a small amount of Blu-Tack adhesive putty or electronics cleaning gel to pull out grime. Throw used putty in trash. If there are tight areas that cannot be reached with wipe, wrap a wipe over the tip of a cotton swab to scrub these areas. Throw wipes and cotton swabs used for cleaning into the trash.

  - For microphones that have a soft windscreen cover (e.g. ChatterVox), remove and throw away. New, fresh windscreen cover will be applied prior to distribution to consumer.

  - If soft surface cases or packaging are visibly dirty, clean using shop towel with dish soap and water followed by wipe with just water to remove soap. Allow to completely air dry before disinfection.

- Disinfect:

  - Wipe the entire device including the screen, covers, accessories, packaging and all cords with a hard surface disinfectant wipe. Reapply disinfectant product as needed to achieve required contact time.

  - Spray soft cases and packaging with soft surface disinfectant. Reapply as needed to achieve required contact time. Allow to air dry prior to repacking equipment.

- Wipe down screens to prevent discoloring:

  - For electronic AAC that have a screen, wipe down the viewing area with PDI Easy Screen Cleaning Wipes or other screen cleaning wipe after the disinfectant wipe has dried to prevent discoloration.

- Put cleaned and disinfected item on “clean” cart to await re-shelving (see last step).

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

How to Clean your Tobii Device

How to Clean PRC/Saltillo Device

How to Clean your AbleNet Device

*(Paper inserts must be laminated to be disinfected)

# Switches and Switch Accessories: Cleaning & Disinfection Procedures

## Examples of Equipment

Switches

Switch Accessible Toys and Accessories

Powerlink

Switch Interfaces

Fabric and Plush Toys Cannot be Loaned Out at this Time

## Supplies Needed

- Blu-Tack adhesive putty, electronics cleaning gel or similar product

- Hard surface disinfectant

- Cotton swabs

- Paper towels or shop towels

- Soft surface disinfectant

- Basin, dish soap and shop towels or sponges

- PPE: gloves and mask

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Unplug all external power sources. Remove batteries for devices with removable batteries.

- Open up any hinged parts and disassemble anything designed to be taken apart.

- Clean prior to disinfecting:

  - Clean hard, non-porous surfaces with a disposable wipe or damp shop towel to remove any chunks, visible dirt and grime on the entire device (including the screen). If there are tight areas that cannot be reached with wipe, wrap a wipe over the tip of a cotton swab to scrub these areas. Throw wipes used for cleaning into the trash.

  - Pillow Switch cover should be removed. Handwash with detergent and water or placed in a mesh laundry bag and wash in washing machine. Air dry completely prior to disinfection.

  - For switches that go in user’s mouth (e.g. sip ‘n puff, etc.), the disposable mouth piece and spit filter should be removed and thrown away. A new one will be sent for the next loan.

  - If soft cases and packaging are visibly dirty, clean using shop towel with dish soap and water followed by wipe with just water to remove soap. Allow to completely air dry before disinfection.

- Disinfect:

  - Apply hard surface disinfectant to all hard, non-porous surfaces including the screen, covers, accessories, and all cords. Reapply as needed to achieve the required contact time.

  - Spray pillow switch cover, soft cases and packaging with soft surface disinfectant. Spray again to achieve the required contact time.

- All toys must be rinsed with water after disinfection to remove any residue as there is a chance that toddlers and young children can mouth these items. Rinsing can be completed by wiping with a wet (water only) paper towel or shop towel.

- Put cleaned and disinfected item on “clean” cart to await re-shelving (see last step).

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

How to Clean your AbleNet Device

Cleaning Toys in Healthcare Blog by CloroxPro

Toy Selection, Cleaning and Sanitation Guidelines from Infection Prevention in Alternative Care Settings

# Equipment for Eating and Food Preparation: Cleaning & Disinfection Procedures

## Examples of Equipment

Anything that touches food or drink

Cutting boards

Liftware and other adapted spoons, forks & knives

Obi Feeding Robot

Cups & mugs

Measuring cups and spoons

## Supplies Needed

- Dish soap

- Disposable sponges or shop towels

- Dish washing basin and dish drying rack

- Food contact surface sanitizer

- Bleach and chlorine test strips (if using bleach rinse as alternative)

- Hard surface disinfectant

- Soft Surface disinfectant

- PPE: gloves and mask

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Sinks, basins, drying rack, or other receptacles used for cleaning equipment and utensils must be cleaned and sanitized before use.

  - Scrub sink, basin and drying rack with disposable sponge and dish soap. Rinse soap off.

  - Spray sink, basin and drying rack with food contact surface disinfectant. Reapply as needed to achieve required contact time. Alternative hard surface disinfectants may be used but MUST be rinsed after disinfection is achieved and prior to use with food contact items.

- Open up any hinged parts and disassemble anything designed to be taken apart (e.g. cutting board attachments).

- Clean prior to disinfecting:

  - Items that are not electric:

    - Pre-rinse and scrape food chunks off item.

    - Wash item with hot water: Scrub with disposable sponge and dishwashing soap to remove any visible particles and grime.

    - Discard used wash water in drain.

  - Electronic items such as Obi Feeding Robot and Liftware: Remove the spoon, plate and placemat and handwash as directed above for items that are not electric. For electronic parts that DO NOT go in users mouth or touch food:  Clean the Obi body and arm and Liftware computerized handle and all cords by wiping damp shop towel or sponge. The Obi body and Liftware handle CANNOT be submerged in water.

  - Spot clean soft cases and packaging using a shop towel with dish soap and water followed by a shop towel with just water to remove soap. Allow to completely air dry before disinfection.

- Disinfect:

  - Items that are not electronic: Spray with Purell Foodservice Surface Sanitizer spray and ensure surfaces stay wet required contact time.

    - If bleach rinse process is substituted follow the specific instructions listed below for dilution and contact time.

  - Electronic items (e.g. Obi, Liftware): The spoon, plate and placemat should be disinfected as directed above for items that are not electric.  For electronic parts that DO NOT go in users mouth or touch food:  apply hard surface disinfectant to the entire device including the screen, covers, accessories, and all cords. Reapply as needed to achieve required contact time.

  - Spray soft cases and packaging with soft surface disinfectant. Spray again as needed to achieve the required contact time.

- Put cleaned and disinfected item on “clean” cart to await re-shelving (see last step).

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## Warnings

- The proper disinfectant must be used for each part. PDI Sani-Cloth AF3 Germicidal Disposable Wipe should NOT be used on food preparation surfaces. Purell Foodservice spray should NOT be sprayed onto computerized parts.

- Computerized parts CANNOT be submerged.

## References

Montana Rule 37.111.321 FOOD EQUIPMENT AND UTENSILS: CLEANING AND SANITIZATION (only regs found that apply to handwashing)

How to Clean Obi Robot and Liftware

## Bleach Rinse as Alternative

- Diluted household bleach solutions may also be used if appropriate for the surface.

  - Check the label to see if bleach is intended for disinfection, and ensure the product is not past its expiration date.

- To make a bleach solution:

  - Mix 1 tablespoon bleach per gallon of water.

  - Determine dilution strength with chlorine ppm test strips – 200 ppm required.

- Follow manufacturer’s instructions for application and proper ventilation. Never mix household bleach with ammonia or any other cleanser.

- Leave solution on the surface for at least 2 minutes.

- Place items in drying rack, allow to air dry.

- Bleach solutions will be effective for disinfection up to 24 hours.

# Telepresence Robots: Cleaning & Disinfection Procedures

## Examples of Equipment

VGo Robots

Ohmni Robot

Kubi Robot

## Supplies Needed

- Disposable, scratch-free, lint-free wipes such as KimTech Delicate Task Wipers

- Screen and lens disinfectant

- Soft surface disinfectant

- Basin, dish soap and shop towels or sponges

- PPE: gloves and mask

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Unplug all external power sources.

- Open up any hinged parts (e.g. camera on VGo) and disassemble anything designed to be taken apart.

- Clean prior to disinfecting:

  - Slightly dampen a KimTech Delicate Task Wiper to remove any visible dirt and grime. If visible dirt is not removed with dampened wipe, try cleaning with an alcohol prep pad. Skip this step for the camera lens, display screen and touch control panel. Do include wheels, charging base and any accessories.

  - If speakers are visibly dirty, use a small amount of Blu-Tack adhesive putty or electronics cleaning gel to pull out grime. Throw used putty in trash.

  - Spot clean soft cases and packaging using shop towels with dish soap and water followed by wipe with just water to remove soap. Allow to completely air dry before disinfection.

- Disinfect:

  - Use a PDI Easy Screen Cleaning Wipe to gently wipe the display screen, camera lens, and touch control panel.

  - Use a 70% isopropyl alcohol prep pad to gently wipe the rest of entire device including the charging base, cords, and accessories.

  - Spray soft cases and packaging with soft surface disinfectant. Reapply as needed to achieve required contact time.

- Put cleaned and disinfected item on “clean” cart to await re-shelving (see last step).

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## Warnings

- Don't get moisture into any openings. All wipes should be slightly damp, not wet.

- Don't spray any cleaners directly onto the item.

## References

VGo Cleaning Guide v.1.1 (saved with MonTECH Cleaning Guide), Step 1 Everyday Cleaning and using disinfectant for hospital settings.  If robot sensors are not working properly or robot is very dirty, perform a deep clean following the directions in VGo guide, but replace the following recommendations:

- Instead of Windex, use 70% isopropyl alcohol wipes and PDI Easy Screen Cleaning Wipes

- Instead of compressed air which can make dust particles with germs (like COVID-19) airborne, use an electronics/keyboard vacuum or soft, lint-free dry wipes like KimTech. If vacuum is used, disinfect vacuum head after use with soft surface disinfectant.

Ohmni and Kubi do not have cleaning guides, so VGo guide will be used for all robots.

# Headsets, Microphones, and FM Systems: Cleaning & Disinfection Procedures

## Examples of Equipment

FM Systems (Contego, Pocket Talker, etc.)

Headsets, Microphones and Speakers (Light Speed, etc.)

## Supplies Needed

- Disposable, scratch-free, soft, lint-free wipes such as KimTech Delicate Task Wipers

- Touch screen disinfectant

- Soft surface disinfectant

- Basin, dish soap and shop towels or sponges

- Tweezers

- PPE: gloves and mask

## Warnings

- Use only a soft, lint-free cloth. Avoid abrasive cloths, towels, paper towels, or similar items

- Don't get moisture into any openings. Wipes should be slightly damp, not wet or dripping.

- Don't spray any cleaners directly onto the item. Always apply cleaner to cloth.

- Do not use bleach, ammonia or abrasive disinfectants.

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Unplug all external power sources and remove batteries, if possible.

- Open up any hinged parts and disassemble anything designed to be taken apart (e.g. detachable microphones, etc.)

- Clean prior to disinfecting:

  - Remove foam headphone ear cushions and immerse in mild soapy water. Set aside and allow to air dry.

  - Remove and discard microphone wind screens. Fresh, new windscreen will be applied prior to distribution to consumer.

  - If visible debris in headset speaker or microphone mesh, very gently clean with dry cotton swab or gently use tweezers to dislodge particles. Do not use liquid on these parts. Do not vacuum or use compressed air.

  - If visible dirt and grime, use a dry KimTech Delicate Task Wipers (or PDI Easy Screen Cleaning wipe) and gently wipe outside of device, non-removable outside earcups (e.g. Bose), headband, cords, etc.

  - Spot clean soft cases and packaging using a shop towel with dish soap and water followed by wipe with just water to remove soap. Allow to completely air dry before disinfection.

- Disinfect:

  - Use a PDI Easy Screen Cleaning Wipe to gently wipe the body of the device (avoiding headset speaker or microphone mesh), including all cords.

  - After the removable foam ear cushions have dried completely spray with soft surface disinfectant and allow surface to remain damp for required contact time.

    - Ear cushions may be replaced if in poor condition or single-use replacements may be supplied if item does not require specific headset.

  - Spray soft cases and packaging with soft surface disinfectant. Surface must remain damp for the required contact time. Reapply as needed.

- Place cleaned and disinfected items on “clean” cart to await re-shelving.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Use clean cart to transport items to storage and re-shelve.

## References

How to clean Pocket Talker

How to clean Contego: Does not recommend cleaning with anything except dry cloth. However, this will not clean nor disinfect the device, so the gentlest cleaning and disinfection procedures are used.

Headsets: Bose (similar model), AfterShokz

- AfterShokz recommends only dry cloth, which will not clean nor disinfect.

- Bose recommends only mild soap and water, which will not disinfect.

How to clean ear buds Bose and Apple AirPods

# Rifton Pacer Gait Trainers, Creepster Crawler: Cleaning & Disinfection Procedures

Examples of Equipment: 8889a, 0270, 1093, 0265B, 0267

## Manufacturer’s Instructions

Rifton: As needed, clean with disinfectant wipes or a solution of up to 10% bleach. Do not use excessive amounts of water. The straps with hook and loop closures may be laundered. Engage the closures before washing. Do not iron. Wash casters with water after outdoor use. Avoid mud and sand.

Creepster Crawler: None.

## Supplies Needed

- PPE: gloves and mask

- basin of soap (dish soap) and water

- disposable shop towels or sponges

- Pick, tweezers and tools as needed for wheels

- Hard surface disinfectant

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Open any hinged parts and remove any pieces intended to detach.

- Clean prior to disinfecting:

  - If straps, harness, or positioner items can be laundered place in mesh laundry bag and wash then air dry. If soft surface items cannot be laundered spot clean with a shop towel and soapy water. Rinse with clean water and a new shop towel to remove soap residue. Allow to air dry prior to disinfection.

  - Remove debris from wheels. Detach wheels if necessary to remove hair, string or other soil.

  - Wet a shop towel in the basin of soap and water, damp enough to clean item and wheels but not dripping water. Lightly scrub all surfaces of the item with the wet shop towel, including the wheels, replace shop towel with a new one when it is dirty. Allow frame to air dry.

  - Discard soapy water and place all towels and sponges in the trash.

- Disinfect:

  - Wipe down all non-fabric surfaces of the item with hard surface disinfectant. Reapply as needed to achieve required contact time.

  - Spray straps, harness and positioner items with disinfecting spray. Soft surfaces must remain damp for required contact time. Reapply spray as needed then allow to air dry.

- Reattach wheels as necessary.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

- Rifton Mini Pacer Gait Trainer Manual: Cleaning p. 26

  - https://cdn.rifton.com/-/media/files/rifton/product-manuals/mini-pacer-gm80.pdf?d=20200618T142706Z

- Rifton Mini and Small Pacer Gait Trainer Manual: Cleaning p. 26

  - https://cdn.rifton.com/-/media/files/rifton/product-manuals/pacers-small-mini-na49.pdf?d=20200623T115414Z

- Rifton Medium, Large and XL Pacer Gait Trainer Manual: Cleaning p. 30

  - https://cdn.rifton.com/-/media/files/rifton/product-manuals/pacers-medium-large-xl-ml34.pdf?d=20200617T162217Z

- Rifton New Pacer Gait Trainer Manual: Cleaning p. 34

  - https://cdn.rifton.com/-/media/files/rifton/product-manuals/dynamic-pacer-pv58.pdf?d=20200218T144644Z

# Crocodile, Pony, and Mustang Gait Trainers: Cleaning & Disinfection Procedures

Examples of Equipment: 0634, 0143, 0728, 0216, 0339, 8765

## Manufacturer’s Instructions

Crocodile Gait Trainer - Cleaning by material type:

Powder coated parts, Stainless steel, Plastics and PU: Wipe of any residue and dirt using a cloth wrung in a pH neutral mix of warm water and non-abrasive cleaning agent. Finish off with a dry lint free cloth to remove what remains of the mix from the surfaces. Disinfect parts when required, pay special attention to any surface that comes into contact with the user/caregiver. Cushions and covers can be cleaned according to the washing label placed upon those. If no washing label is attached, then we do not recommend washing the item

Upholstery: Do not wash upholstery. Replace upholstery when required Ware-washing machine A wide selection of our products can be cleaned in a Ware-washing machine at either 60° or 85°C. Please refer to the product User Guide for detailed information.

Cleaning agents: Use a mild non-abrasive cleaning agent with a PH value between 5-9. The cleaning agent must be used in a mix with water. The mix of the cleaning agent and water must be pH neutral. Do not use Acidic or Alkaline cleaning agents directly on any surface as they may damage the surface

Disinfectant: Surfaces can be disinfected with a 70% IPA solution

Hygienic measures:

What: All products from R82

Who: Distributors, owners, caregivers and authorized service staff

When: As required, scheduled maintenance, at repair or before reuse

Measures: We recommend wearing protective gloves when cleaning. Do not let water ingress into cables, bearings or frame parts. Any actuator should be retracted to minimum stroke distance before cleaning

How: First clean the surfaces then wipe the surfaces dry with a lint free cloth or towel. Subsequently an air dryer at low temperature can be used

Disinfection: It is recommended to wipe of any residue and dirt using a cloth wrung in a pH neutral mix of warm water and non-abrasive cleaning agent and let the surfaces dry before disinfection. Afterwards the surfaces can be disinfected with a 70% IPA solution

Pony Walker: The Pony functions without any maintenance, but it is recommended to clean the frame frequently. Do not use any cleaning materials containing chlorine.

Mustang Walker: The product without accessories can be disinfected with a 70% disinfectant IPA solution. It is recommended to wipe off any residue and dirt from the product, using a cloth with warm water and a mild detergent or soap without chlorine and let it dry before disinfection.

## Supplies Needed

- PPE: gloves and mask

- basin of soap (dish soap) and water

- disposable shop towels or sponges (lint-free)

- tweezers, pick and tools as needed for wheels

- Alcohol prep pads, liquid 70% isopropyl alcohol or PDI Easy Screen Cleaning Wipes

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Open any hinged parts and remove any pieces intended to detach.

- Clean prior to disinfecting:

  - Remove cloth cover and foam bar pad. Hand wash or launder cloth cover as appropriate. Wipe down foam pad with soapy water and sponge or shop towel. Allow all soft surface items to air dry prior to disinfecting.

  - Remove debris from wheels. Detach wheels if necessary to remove hair, string or other soil.

  - Wet a shop towel in the basin of soap and water, damp enough to clean gait trainer and wheels but not dripping water.

  - Lightly scrub all surfaces of the gait trainer with the wet shop towel, including the wheels, replace shop towel with a new one when it is dirty. Allow frame to air dry.

  - Discard soapy water and used shop towels.

- Disinfect:

  - Wipe down all surfaces of the gait trainer with an alcohol prep pad or 70% isopropyl alcohol applied to a shop towel.

  - Spray cloth cover and foam pad with disinfecting spray. Surface must remain damp for required contact time. Reapply spray as needed. Allow to air dry prior to reattaching.

- Reattach wheels as necessary.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

- R82 Crocodile User Guide: Disinfection p. 31

  - http://www.r82.com/pimmedia/document/m1220_crocodile_8_48326.pdf

- R82 Pony User Manual: Maintenance p. 4

  - http://www.r82.com/pimmedia/document/pony_en_48588.pdf

- R82 Mustang User Guide: Disinfection p. 28

  - http://www.r82.com/pimmedia/document/m1250_mustang_9_48483.pdf

# Rifton TRAM: Cleaning & Disinfection Procedures

Examples of Equipment: 0281, 0738

## Manufacturer’s Instructions

To minimize risk of infectious disease transmission, clean and disinfect the TRAM between uses by different patients. To avoid damaging the product:

• Do not use excessive amounts of water when cleaning the TRAM.

• Do not use high-pressure spray or steam cleaning.

• Do not clean the piston rod of the TRAM’s electric actuator.

The TRAM and its accessories may be cleaned with broad spectrum, multi-purpose disinfectant sprays and wipes or a solution of up to 10% bleach. Do not use solutions containing perfumes or staining ingredients. The leg straps, pelvic support and removable back belt (if installed) may be machine washed in cold water and air dried.

## Supplies Needed

- PPE: gloves and mask

- basin of soap (dish soap) and water

- disposable shop towels or sponges

- Hard surface disinfectant

- laundry detergent

- Soft surface disinfectant

- washing machine

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Clean prior to disinfecting:

  - Remove leg straps, pelvic support and back belt. Place items in mesh laundry bag and machine wash. Allow to air dry prior to disinfection.

    - If soft surface (straps, belts, etc.) items cannot be removed spot clean prior to disinfection using a shop towel dipped in soapy water. Use a shop towel and clean water to rinse soap residue.

  - Remove debris from wheels. Detach wheels if necessary to remove hair, string or other soil.

  - Wet a shop towel in the basin of soapy water, damp enough to clean TRAM and wheels but not dripping water.

  - Lightly scrub all surfaces of the TRAM with the wet shop towel, including the wheels. Replace shop towel with a new one when it is dirty. Do not clean the piston rod of the TRAM’s electric actuator. Allow frame to air dry

  - Discard soapy water and used shop towels.

- Disinfect:

  - Wipe down all surfaces of the pacer gait trainer with a hard surface disinfectant. Surfaces must remain wet for required contact time. Allow surfaces to air dry then reassemble as needed.

  - Spray soft surface items (straps, positioners, harnesses, belts, etc.) with disinfecting spray. Surfaces must remain damp for required contact time. Reapply spray as needed. Allow to air dry and reattach.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Return item to designated storage location

## References

- Rifton TRAM Manual: Cleaning p. 26

  - https://cdn.rifton.com/-/media/files/rifton/product-manuals/rifton-tram-hl99.pdf?d=20190821T150220Z

# Tumble Forms 2 Products, Special Tomato Soft Touch Floor Sitter with Material or Wooden Wedge: Cleaning & Disinfection Procedures

Examples of Equipment: 0553, 0454, 0340, 0342, 0644, 0587, 0590, 9409, 0348, 0345, 0497

## Manufacturer’s Instructions

Tumble Forms 2: Cleans easily with damp cloth and common disinfectants. Tumble Forms 2® covering may be cleaned with household grade detergents, either liquid or abrasive powder types. Permanent markers and ballpoint pen markings cannot be removed. Non-permanent inks, wax, crayon, oil, grease or dirt can be scrubbed with an abrasive powder cleanser or scouring pad and wiped clean with a damp cloth. (Do not soak the cushion, as it will get water-logged).

Special Tomato Products are impermeable to fluids, but do not submerge. Use a mild detergent/cleaner to wipe the chair periodically. To extend the life of the hook and loop strips, comb out the hook portion (hard side) to remove lint and other debris. HARNESS CARE: Place harness in a mesh bag and machine wash with cold water, then air dry.

## Warnings

Do not allow water to enter vent. Do not submerge or soak any items other than straps, belts or harnesses.

## Supplies Needed

- PPE: gloves and mask

- Basin of soapy water

- Disposable shop towels or sponges

- Hard surface disinfectant

- Soft surface disinfectant

- Laundry detergent

- Washing machine

- Fine tooth comb and vacuum as needed

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Clean prior to disinfecting:

  - Remove harness and straps, place in mesh laundry bag and machine wash. Allow to air dry.

  - Wet a shop towel or sponge in the basin of soapy water, damp enough to clean seat and wedge but not dripping water. Lightly scrub all surfaces of the seat and wedge with the wet shop towel, replace shop towel with a new one when it is dirty.

  - Use fine tooth comb and vacuum as needed to remove debris from Velcro.

  - Allow seat and wedge to air dry.

  - Discard soapy water and used shop towels.

- Disinfect:

- Wipe down all surfaces of the seat and wedge with hard surface disinfectant while maintaining required contact time.

- Spray vent, Velcro, harness and straps with soft surface disinfectant. Reapply spray as needed to achieve required contact time. Allow to air dry prior to reattaching.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

- Special Tomato Soft-Touch Sitter (0587) User Manual: Harness Diagram p. 2, Cleaning p. 6

  - https://www.specialtomato.com/media/wysiwyg/PDFs/sitters.pdf

- Special Tomato MPS Hi-Low Seating System (1075) Manual: Harness Threading p. 14, Cleaning p. 15

  - https://www.specialtomato.com/media/wysiwyg/PDFs/hilow.pdf

- Special Tomato Soft Touch Booster – Discontinued (1109): Product Care p. 2

  - https://www.specialtomato.com/media/wysiwyg/PDFs/boosters.pdf

- Tumble Forms 2 Guide: Care Information p. 19

  - http://www.rehabandmedical.co.uk/pdf/Tumble-Forms.pdf

- Keekaroo Café Booster Manual: Product Care p. 2

  - https://www.keekaroo.com/media/wysiwyg/Manuals/Cafe_Booster_UM_English_9.11.13.pdf

# Special Tomato Height Right, Keekaroo Height Right and Stokke Tripp Trapp Chairs (wooden with straps): Cleaning & Disinfection Procedures

Examples: 0065, 0024, 0445, 0474

## Manufacturer’s Instructions

Stokke Tripp Trapp: Wipe with a clean damp cloth, wipe off excess of water with a dry cloth. Moisture will leave cracks in stain. We do not recommend the use of any detergent. Colors may change if the high chair is exposed to sunlight.

Special Tomato Height Right: Clean chair with damp cloth; then wipe dry.

Keekaroo Height Right:

## Warnings

Take note of wood finish condition. If finish is worn and wood is no longer sealed the item will need to be set aside for repair after cleaning.

## Supplies Needed

- PPE: gloves and mask

- Basin of soapy water

- Disposable shop towels or sponges

- Hard surface disinfectant

- Soft surface disinfectant

- Washing machine and laundry detergent

## Cleaning & Disinfection Instructions:

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Detach all pieces that are intended to come apart.

- Clean prior to disinfecting:

- Remove straps, place in mesh laundry bag and machine or hand wash. If straps can’t be removed spot clean with a shop towel and soapy water. Rinse with clean water and a new shop towel to remove soap residue. Allow to air dry.

- Wet a shop towel or sponge in the basin of soapy water, damp enough to clean seat but not dripping water. Lightly scrub all surfaces of the seat with the wet shop towel, replace shop towel with a new one when it is dirty. Allow chair to air dry.

- Discard soapy water and used shop towels.

- Disinfect:

- Wipe down all surfaces of the chair with hard surface disinfectant while maintaining required contact time.

- Spray straps with soft surface disinfectant. Reapply spray as needed to achieve required contact time. Allow to air dry prior to reattaching.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

- Special Tomato Height Right Chair Manual – Discontinued (0573): Product Care p. 7

  - https://www.specialtomato.com/media/wysiwyg/PDFs/heightright.pdf

- Special Tomato Comfort Cushions – Discontinued (6679A): Product Care p. 2

  - https://www.specialtomato.com/media/wysiwyg/PDFs/heightrightcc.pdf

- Special Tomato Soft Touch Booster – Discontinued (1109): Product Care p. 2

  - https://www.specialtomato.com/media/wysiwyg/PDFs/boosters.pdf

- Stokke Tripp Trapp Chair Manual: Cleaning and maintenance p. 6, Harness diagram p. 16

  - https://www.stokke.com/on/demandware.static/-/Sites-stokke-master-catalog/default/dw7ed61bec/documents/Tripp Trapp user guide US - WEB.pdf

- Keekaroo Height Right Chair Manual: Product Care p. 6

  - https://www.keekaroo.com/media/wysiwyg/Manuals/HRHC_II_Convert_UM_-_2016-09_-_Low_Res_WEB_1.pdf

- Keekaroo Café Booster Manual: Product Care p. 2

  - https://www.keekaroo.com/media/wysiwyg/Manuals/Cafe_Booster_UM_English_9.11.13.pdf

# Kaye Tilting Benches, Classroom Positioner Chair: Cleaning & Disinfection Procedures

Examples: 0665, 0405, 0611, 8071a

## Manufacturer’s Instructions

Kaye: Do not expose the unit to rain or submerge it in water. Use a slightly damp cloth to clean the unit, but wipe the dampness from the wood after cleaning it.

CPC: No instructions

## Warnings

Take note of soft surface, vinyl cover and wood finish condition. If cover is torn or has holes the item must be removed from loan pool (after cleaning) until it has been repaired. If wood finish is worn and wood is no longer sealed the item will need to be repaired (after cleaning) prior to any further loan.

## Supplies Needed

- PPE: gloves and mask

- Basin of soapy water

- Disposable shop towels or sponges

- Small scrub brush or toothbrush

- Hard surface disinfectant

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Detach all pieces that are intended to come apart.

- Clean prior to disinfecting:

- Remove any straps or fabric pieces that can be removed and hand or machine wash. If fabric items can’t be removed, spot clean using a shop towel and soapy water. Rinse with clean water and a new shop towel to remove soap residue. Allow to air dry prior to disinfection.

- Clean hair, dirt and debris from wheels. Remove wheels if necessary.

- Wet a shop towel or sponge in the basin of soapy water, damp enough to clean item but not dripping water. Lightly scrub all surfaces of the item, including wheels, with the wet shop towel. Replace shop towel with a new one when it is dirty. Gently clean rubber gripping “feet” with a damp scrub brush or toothbrush as needed. Allow item to air dry.

- Discard soapy water and used shop towels.

- Disinfect:

- Spray gripping “feet”, straps and fabric pieces with disinfecting spray. Reapply as needed to achieve required contact time. Allow to air dry prior to reattaching.

- Wipe down all other surfaces with hard surface disinfectant while maintaining required contact time. Allow item to dry completely. Reattach wheels if needed.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

- Kaye Adjustable Benches Instructions: Maintenance and Care p. 2

  - kayeproducts.com > Consult Us > Product Instructions > Kaye Adjustable Benches

- Prospect Designs Classroom Positioner Chair User Manual: No Cleaning Instructions

  - prospectdesigns.com > Products > Classroom Positioner Chair > CPC User Manual

# Wenzelite Chair, Prime Engineering Stander and Mulholland Standers:

Examples: 0566, 0567, 0478, 0499, 9499a, 0144, 0477

## Manufacturer’s Instructions:

Wenzelite: The device is cleaned down with a damp cloth using mild detergent and/or disinfectant. For example, a mild chloride disinfectant (solution of 1000 parts per million) can be used. Excess water is then removed from the device using a towel or a compressed air tool. The device is then placed in a suitable environment to fully dry.

Prime: Use soft cloth when cleaning table top. All fabrics should be hand washed with mild soap and water, then allowed to air dry.

Mulholland: Manufacturer is no longer in business and no instructions are available.

## Supplies Needed

- PPE: gloves and mask

- Basin of soapy water

- Disposable shop towels or sponges

- Hard surface disinfectant

- Soft surface disinfectant

- Screwdriver, Allen wrench and assorted tools

- Washing machine and laundry detergent

- Fine tooth comb and vacuum as needed

## Cleaning & Disinfection Instructions:

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Open all hinges or points of attachment and remove all pieces that detach.

- Clean prior to disinfecting:

- Use fine tooth comb and vacuum as needed to remove debris from Velcro. Remove harness and straps, place in mesh laundry bag and machine wash. If fabric items can’t be removed, spot clean with a shop towel and soapy water. Rinse with clean water and a new shop towel to remove any soap residue. Allow to air dry.

- Remove dirt, hair and debris from wheels. Remove wheels if necessary.

- Wet a shop towel or sponge in the basin of soapy water, damp enough to clean frame surfaces and trays but not dripping water. Lightly scrub all surfaces of the item, including wheels, with the wet shop towel. Replace shop towel with a new one when it is dirty. Allow frame and tray to air dry.

- Discard soapy water and used shop towels.

- Disinfect:

- Wipe down all surfaces of the frame and tray, including wheels, with hard surface disinfectant while maintaining required contact time.

- Spray harness and straps with soft surface disinfectant. Reapply spray as needed to achieve required contact time. Allow to air dry prior to reattaching.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

- Inspired by Drive First Class Chair (0566) User Manual: Cleaning Procedure

  - https://www.inspiredbydrive.com/wp-content/uploads/2016/11/FC_2000N_FC-4000N_hangtag.pdf

- Prime Engineering Superstand HLT Manual: Cleaning and Care p. 5

  - https://primeengineering.com/wp-content/uploads/2018/01/Superstand-HLT-Owners-Manual.pdf

- Prime Engineering Superstand Youth Manual: General Maintenance p. 14

  - https://primeengineering.com/wp-content/uploads/2018/01/Superstand-Youth-Owners-Manual.pdf

# Rifton Supine and Prone Stander: Cleaning & Disinfection Procedures

Examples: 9401, 9402, 0417

## Manufacturer’s Instructions

Clean with disinfectant wipes or a solution of up to 10% bleach. Do not use excessive water on pads and wooden surfaces. Clean tray with mild soap or detergent. Avoid abrasive cleaners and aromatic spirits. Disinfectants can be used, but no solvents. Never scrape with razor blades or other sharp instruments. Do not clean in hot sun or at elevated temperatures. Do not send through a dishwasher. The components of the arm Anchor can be removed for cleaning. Wipe clean with disinfectant wipes or a solution of up to 10% bleach.

## Warnings

Take note of soft surface, vinyl cover and wood finish condition. If cover is torn or has holes the item must be removed from loan pool (after cleaning) until it has been repaired. If wood finish is worn and wood is no longer sealed the item will need to be repaired (after cleaning) prior to any further loan.

## Supplies Needed

- PPE: gloves and mask

- Basin of soapy water

- Disposable shop towels or sponges

- Small scrub brush or toothbrush

- Hard surface disinfectant

- Soft surface disinfectant

- Washing machine and detergent

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Open all hinges or points of attachment. Remove all pieces that are intended to come apart.

- Clean prior to disinfecting:

- Remove any straps or fabric pieces that can be removed and hand or machine wash. If fabric items can’t be removed, spot clean using a shop towel and soapy water. Rinse with clean water and a new shop towel to remove soap residue. Allow to air dry prior to disinfection.

- Clean hair, dirt and debris from wheels. Remove wheels if necessary.

- Wet a shop towel or sponge in the basin of soapy water, damp enough to clean item but not dripping water. Lightly scrub all surfaces of the item, including wheels, with the wet shop towel. Replace shop towel with a new one when it is dirty. Allow item to air dry.

- Discard soapy water and used shop towels.

- Disinfect:

- Spray straps and fabric pieces with disinfecting spray. Reapply as needed to achieve required contact time. Allow to air dry prior to reattaching.

- Wipe down all other surfaces with hard surface disinfectant while maintaining required contact time. Allow item to dry completely. Reattach wheels if needed.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

- Rifton Prone Stander (0417) Product Manual: Cleaning p. 18

  - https://cdn.rifton.com/-/media/files/rifton/product-manuals/prone-stander-vh97.pdf?d=20180426T185529Z

- Rifton Supine Stander (9401) Product Manual: Cleaning p. 18

  - https://cdn.rifton.com/-/media/files/rifton/product-manuals/supine-standers-kw26.pdf?d=20200618T190126Z

# EasyStand Standing Frames: Cleaning & Disinfection Procedures

Examples: Evolv 3968a and Bantam 8190a, 0582

## Manufacturer’s Instructions

Clean the EasyStand with a mild, non-abrasive disinfectant or soap and water.

• Do not use bleach or any other bleach-based cleaners on the EasyStand.

• Disinfecting wipes are safe to use on metal and non-porous surfaces.

• Use a damp sponge with mild soap or a mild, non-abrasive disinfectant.

• Spot clean Neoprene with mild soap and water.

• Neoprene covers may be removed and hand washed with water. Air dry the covers. Always use care when removing and replacing the covers.

• Do not dry clean any upholstery items. The recommended cleaning and disinfection instructions should be followed whenever possible to avoid damage to the PU membrane which could lead to premature failure of the cover.

• For superficial dirt, use a disposable wipe and a warm solution of neutral detergent. Do not use abrasive cleaners.

• Disinfect using a 0.1% sodium hypochlorite solution (1000ppm chlorine), stronger bleach solutions not recommended.

• Thoroughly rinse any remaining cleaning solution with water and dry fabric before re-use or storage.

• Do not iron

## Supplies Needed

- PPE: gloves and mask

- Basin of soap (dish soap) and water

- Disposable shop towels or sponges

- Hard surface disinfectant

- Laundry detergent

- Soft surface disinfectant

- Washing machine

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Clean prior to disinfecting

  - Remove neoprene covers and spot clean or hand wash in basin of soapy water. Allow to air dry.

  - Remove all other soft surface items, place in mesh laundry bag and machine wash. Allow to air dry prior to disinfecting.

  - Clean hair, dirt and debris from wheels. Remove wheels if necessary.

  - Wet a shop towel or sponge in the basin of soapy water, damp enough to clean but not dripping water. Lightly scrub all surfaces of the stander with the wet shop towel or sponge, including the wheels, replace shop towel with a new one when it is dirty. Allow frame to air dry.

  - Discard soapy water and used shop towels.

- Disinfect:

  - Wipe down all surfaces of the stander with hard surface disinfectant. Reapply as needed to achieve required contact time.

  - Spray straps, harnesses and covers with soft surface disinfectant. Soft surfaces must remain damp for required contact time. Reapply as needed then allow to air dry prior to reattaching.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Return item to designated storage location.

## References

- EasyStand Evolv (3968A) Manual: Cleaning p. 38

  - https://easystand.com/owners-manuals/Evolv M L XT.pdf

- EasyStand Bantam (8190A) Manual: Cleaning p. 30

  - https://easystand.com/owners-manuals/Bantam XS S.pdf

# Rifton Large Dynamic Stander and TumbleForms Tristander 45:

Examples: 1095, 0047B

## Manufacturer’s Instructions

Rifton: Clean with disinfectant wipes or a solution of up to 10% bleach. Do not use excessive amounts of water. The fleece sleeve can be removed for laundering. Do not iron.

TumbleForms: To clean the Tristander support pads use mild detergent on gently cycle. Tumble dry. The Tumble Forms covering may be cleaned with household or industrial grade detergents, either liquid or abrasive powder types. Do not soak or immerse the cushion in water. The tumble forms covering can be disinfected with alcohol or 3% hydrogen peroxide. It can be scrubbed clean with a nylon brush.

## Warnings

Take note of the vent in each TumbleForms support section. Do not submerge or allow fluids to enter vent other than soft surface disinfectant.

## Supplies Needed

- PPE: gloves and mask

- Basin of soap (dish soap) and water

- Disposable shop towels or sponges

- Hard surface disinfectant

- Washing machine and laundry detergent

- Soft surface disinfectant

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Open any hinged parts and remove any pieces intended to detach.

- Clean prior to disinfecting:

  - Remove fleece sleeve, straps and fabric pieces then place in mesh laundry bag. Machine wash. Allow to air dry prior to disinfecting.

  - Remove dirt, hair and debris from wheels. Remove wheels if necessary to clean.

  - Wet a shop towel or sponge in the basin of soap and water, damp enough to clean but not dripping water. Lightly scrub all surfaces with the wet shop towel or sponge, including the wheels. Replace shop towel or sponge with a new one when it is dirty. Allow frame to air dry

  - Discard soapy water and used shop towels.

- Disinfect:

  - Wipe down all surfaces with hard surface disinfectant and maintain required contact time. Allow frame and support cushions to air dry.

  - Spray sleeve, straps and fabric pieces with soft surface disinfectant. Reapply as necessary to achieve required contact time. Allow to air dry prior to reattaching.

  - Spray TumbleForms vent on each section of cushion with soft surface disinfectant.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Return item to designated storage location.

## References

- Special Tomato MPS Hi-Low Seating System (1075) Manual: Harness Threading p. 14, Cleaning p. 15

  - https://www.specialtomato.com/media/wysiwyg/PDFs/hilow.pdf

- Tumble Forms 2 Guide: Care Information p. 19

  - http://www.rehabandmedical.co.uk/pdf/Tumble-Forms.pdf

- Tumble Forms Tristander 45 (0047B) Manual: Cleaning p. 15

  - https://www.medicaleshop.com/pdfs/tumble-forms/manual/tristander-ins.pdf

- Rifton Dynamic Stander (1095) Manual: Cleaning p. 14

  - https://cdn.rifton.com/-/media/files/rifton/product-manuals/mobile-standers-vj47.pdf?d=20180423T142820Z

# Rifton Pacer Accessories: Cleaning & Disinfection Procedures

Examples: Arm prompts, hand loops, chest prompt, pelvic support, hip positioner

## Manufacturer’s Instructions:

Adjustable arm prompt base and are made of closed-cell molded foam for padded comfort and easy cleaning. Straps are typically made of polypropylene or nylon webbing.

Upholstery items (pads, support blocks, padded prompts, etc.) are typically polyurethane foam with a fire-retardant cover made from expanded vinyl.

## Supplies Needed:

- PPE: gloves and mask

- Basin of soap and water

- Disposable shop towels or sponges

- Hard surface disinfectant

- Soft surface disinfectant

- Washing machine and laundry detergent

- Fine tooth comb, vacuum, swabs and fun-tak as needed

## Cleaning & Disinfection Instructions:

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Open any hinged parts and detach any pieces intended to be removed.

- Clean prior to disinfecting:

  - Remove straps, harness and any other fabric pieces (if removable). Place in mesh laundry bag and machine wash or hand wash. If straps or Velcro are not removable, remove debris with a fine-tooth comb, vacuum, or fun-tak as needed. Spot clean with disposable shop towels and soapy water. Rinse with clean water and a new shop towel to remove any soap residue. Allow to air dry prior to disinfecting.

  - Use a shop towel and soapy water to lightly scrub all surfaces that are metal, plastic, closed-cell molded foam or covered with vinyl. Use cotton swaps to clean crevices if needed. Allow to air dry prior to disinfection.

  - Discard used water and shop towels.

- Disinfect

  - Apply disinfectant spray to all surfaces (non-porous and porous). Reapply as needed to achieve required contact time.

  - Allow all soft surface items to air dry prior to reattaching.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Return item to designated storage location

# Bumbo Baby Seat and Rounded Sit-Me-Up Seat: Cleaning & Disinfection Procedures

Examples: 5909A, 0671, 0670

## Manufacturer’s Instructions

Bumbo: The surface of the product has open cells which will act like a sponge through excessive use of water therefore only use a damp cloth and mild detergent and never immerse the product in water.

Sit-Me-Up: All vinyl and polyurethane covered items can be surface washed with a soft cloth and mild liquid soap and water solution. To disinfect, use up to a 5:1 water/household bleach solution. When cleaning or disinfecting, be sure to thoroughly rinse with a clean wet cloth to remove all chemicals. This will provide the purest surface for your child and also prolong the life of the material.

## Warnings

Take note of cover condition. If cover has rips, tears or holes the item must be removed from inventory after cleaning.

## Supplies Needed

- PPE: gloves and mask

- Basin of soap (dish soap) and water

- Disposable shop towels or sponges

- Soft Surface Disinfectant

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Clean prior to disinfecting:

  - Using a shop towel, dish soap and water (wring out excess water), lightly scrub all surfaces.

  - Discard used water and shop towels.

- Disinfect:

  - Wipe down with food contact surface cleaner or soft surface disinfectant. Reapply as needed to achieve required contact time. Allow to air dry.

  - Use clean water and a new shop towel to rinse all chemical residue from surfaces then allow to air dry.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Place item on “clean” cart and return item to designated storage location.

## References

- Bumbo Floor Seat Product Information: Use & Care

  - http://www.bumbousa.com/product/floor-seat/

- Children’s Factory Resources – cleaning instruction

  - https://www.childrensfactory.com/resources/cleaning-instructions/

# Birillo, LadyBug: Cleaning & Disinfection Procedures

Examples: 0302, 0275

## Manufacturer’s Instructions

Birillo: CLEANING THE PAINTED AND CHROME-PLATED FRAME Clean and disinfect the painted and chromed parts of the seat and base with a soft cloth dipped in a household detergent and disinfectant, such as alcohol diluted in water. CLEANING AND DISINFECTION OF THE CUSHIONS Clean the cushions with a soft cloth dipped in water with detergent and a neutral home disinfectant. CLEANING THE WHEELS At least once a month, clean the wheels of residues (threads, hair, etc.) that can build up on rotating parts. Lubricate sliding parts with spray grease or similar.

LadyBug: No instructions

## Warnings

Note condition of cushion material. If cover is damaged the item must be removed from loan inventory until repaired.

## Supplies Needed

- PPE: gloves and mask

- 70% isopropyl alcohol (liquid applied to shop towel or pre-moistened wipe)

- Hard surface disinfectant

- Tweezers or pick

- Basin and dish soap

- Disposable shop towels or sponges

- Vacuum

- Washing machine and detergent

## Cleaning & Disinfecting Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Open all hinged sections and remove all pieces intended to detach.

- Clean prior to disinfecting:

  - Remove any straps, belts or fabric for machine or hand washing. If fabric items do not detach spot clean with a shop towel and soapy water. Rinse with clean water and new shop towel to remove soap residue. Allow to air dry prior to disinfecting.

  - Vacuum crevices to remove loose dirt. Remove string, hair, dirt and debris from wheels. Remove wheels if necessary.

  - Dip the shop towel or sponge in soapy water then wring out excess water prior to cleaning all surfaces, including wheels and tray. Allow equipment to air dry completely before disinfecting.

  - Discard used water and shop towels.

- Disinfect:

- Wipe metal frame with 70% isopropyl alcohol.

- Wipe cushions and all other surfaces, including wheels, with hard surface disinfectant. Reapply as needed to achieve required contact time.

- Throw used wipes in the trash.

- Clean and disinfect all tools (vacuum head, tweezers, etc.) used to clean each “batch” of equipment or if they become heavily soiled.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves (and any other soiled PPE) and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Return item to assigned storage area.

## References

- Ormesa Birillo Use and Maintenance Handbook: Cleaning & Disinfection p. 19

  - https://mobility-usa.com/pdfs/om-ormesa-birillo.pdf

# Firefly GoTo Seat: Cleaning & Disinfection Procedures

Examples: 9410

## Supplies Needed

- PPE: gloves and mask

- Basin of soapy water

- Disposable shop towels or sponges

- Hard surface disinfectant

- Soft surface disinfectant

- Laundry detergent

- Washing machine

- Fine tooth comb and vacuum as needed

## Cleaning & Disinfection Instructions

- Wash your hands with soap and water.

- Fill up basin with soap and water.

- Put on gloves and mask.

- Open all hinges and detach all pieces that are intended to come apart.

- Clean prior to disinfecting:

- Remove harness, straps, and fabric cover. Use fine tooth comb and vacuum as needed to clean out Velcro fastenings prior to machine washing. Place fabric pieces in mesh laundry bag and machine wash. Allow to air dry.

- Wet a shop towel or sponge in the basin of soapy water, damp enough to clean plastic seat base but not dripping water. Lightly scrub all surfaces of the seat frame and floorsitter base with the wet shop towel, replace shop towel with a new one when it is dirty. Allow frame and base to air dry.

- Discard soapy water and used shop towels.

- Disinfect:

- Wipe down all surfaces of the seat frame and base with hard surface disinfectant while maintaining required contact time.

- Spray straps, harness and fabric cover with soft surface disinfectant. Reapply spray as needed to achieve required contact time. Allow to air dry prior to reattaching.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Using the “clean” cart, transport all items back to storage and re-shelve.

## References

- Firefly GoTo User’s Manual: Cleaning and Care p. 6-7

  - https://www.fireflyfriends.com/media/catalog/product/g/o/goto_user_manual_ls604-04_web.pdf

# Manual Wheelchairs, Walkers, Rollators, Knee Scooters, Canes and Crutches: Cleaning & Disinfecting Procedures

Examples: 8145A, 6985A, 0089B, 0108, 7194A, 6723A, 3658, 0078, 0016, 4768A, 0116, 8095A

## Supplies Needed

- PPE: gloves and mask

- Soft surface disinfectant

- Hard surface disinfectant

- Tweezers or pick

- Basin and soapy water

- Disposable shop towels or sponges

- Vacuum and fine-tooth comb

- Washing machine and detergent

## Cleaning & Disinfecting Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Remove all pieces that detach and fully open all hinged pieces.

- Clean prior to disinfecting:

  - Remove any fabric items that detach and launder. If soft surface items cannot be removed spot clean with shop towel or sponge and soapy water. Rinse with shop towel and clean water. Use basin to hand wash items that are not machine washable. Allow items to air dry prior to disinfecting.

  - String, hair, dirt and debris must be removed from wheels and crutch or cane tips. Remove wheels if necessary. Vacuum crevices, soft surfaces and Velcro to remove loose soil. Apply fine tooth comb to Velcro.

  - Dip a shop towel or sponge in soapy water then wring out excess water prior to cleaning all surfaces. Allow to air dry prior to disinfecting.

  - Throw used towels and sponges in the trash.

- Disinfect:

  - Spray all surfaces (hard and soft surfaces) with disinfectant spray. Reapply as needed to achieve required contact time. Allow all surfaces to dry prior to reattaching or reassembling.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves (and any other soiled PPE) and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Return item to assigned storage area.

## References

- Breezy EC Wheelchair (6985A) Manual: No Cleaning Instructions

  - http://www.sunrisemedical.ca/getattachment/064043d1-f09c-4fd8-9efc-922d9e8c64ee/.aspx

- Sanitizing Assistive Technology, Georgia Tech – Pass It On Center

  - https://pioc.gatech.edu/pioc/kb_article.php?id=61

- Drive Medical Nitro Rollator (1026A) Manual: Cleaning p. 6

  - https://www.drivemedical.com/medias/sys_master/images/images/h13/h76/9459992494110/219862.pdf

# Adaptive Strollers: Cleaning & Disinfecting Procedures

Examples: 9500A, 9361A, 0277

## Supplies Needed

- PPE: gloves and mask

- Hard surface disinfectant

- Soft surface disinfectant

- vacuum

- screwdrivers and wrenches

- tweezers, pick

- disposable sponges or shop towels

- basin of warm soapy water

- user’s guide or instruction manual

## Cleaning & Disinfecting Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Clean prior to disinfecting:

  - Take photos prior to disassembling. Follow manufacturer’s instructions to remove fabric cover and any other soft surfaces to be laundered.

  - Place straps and harnesses in mesh laundry bag and machine wash. Machine or hand wash fabric cover and any other soft surface items following manufacturer’s instructions. If soft surface items cannot be removed spot clean with shop towel and soapy water. Rinse with shop towel and clean water to remove soap residue. Allow to air dry prior to disinfection.

    - Mild, hypoallergenic detergent. Warm water, delicate cycle.

  - Clean any debris and dirt from wheels using tweezer or pick as needed. Remove wheels if necessary.

  - Wipe down frame, wheels and all solid surfaces with disposable sponges or shop towels and warm, soapy water to remove all dirt and debris. Allow to air dry prior to disinfecting.

  - Discard used water and shop towels.

- Disinfect:

  - Wipe or spray all surfaces with disinfectant. Follow label instructions for contact time and reapply as needed. Allow soft surface items to air dry. All surfaces must be completely dry prior to reassembly.

- Wash hands and put on clean PPE prior to reassembly. Refer to photos taken in step 3a. Reattach cover and straps per manufacturer’s instructions.

- Notify appropriate staff of any repairs or maintenance needed (i.e.: wheel lubrication).

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves (and any other soiled PPE) and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Return item to assigned storage area.

## References

- Convaid Basic Cruiser CX16 User’s Guide (9500A): Remove Fabric p. 48, Cleaning p. 57

  - https://www.convaid.com/wp-content/uploads/2019/04/UG0019-ENG-Cruiser-03-16-18.pdf

- Advance Mobility Freedom Push Chair Manual (9361A): Cleaning p. 13

  - https://www.advancemobility.co.uk/manuals/freedom_manual.pdf

- Special Tomato Jogger (0277): Removing Seat Fabric p.4, Cleaning p. 8

  - https://www.specialtomato.com/media/wysiwyg/PDFs/UMJOG06-18LIT-Web.pdf

# AmTryke Early Intervention Tryke: Cleaning & Disinfection Procedures

Examples: 1410

## Manufacturer’s Instructions

No cleaning or disinfection instructions found.

## Supplies Needed

- PPE: gloves and mask

- Soft surface disinfectant

- Hard surface disinfectant

- Tweezers or pick

- Basin and soapy water

- Disposable shop towels or sponges

- Fine-tooth comb

## Cleaning & Disinfecting Instructions

- Wash your hands with soap and water.

- Put on gloves and mask.

- Remove all pieces that detach and fully open all hinged pieces.

- Clean prior to disinfecting:

  - Remove any fabric items that detach and launder. If soft surface items cannot be removed spot clean with shop towel or sponge and soapy water. Rinse with shop towel and clean water. Use basin to hand wash items that are not machine washable. Apply fine-tooth comb to Velcro if necessary to remove hair or debris. Allow items to air dry prior to disinfecting.

  - String, hair, dirt and debris must be removed from wheels. Remove wheels if necessary.

  - Dip a shop towel or sponge in soapy water then wring out excess water prior to cleaning all surfaces, including wheels. Allow to air dry prior to disinfecting.

  - Throw used towels and sponges in the trash.

- Disinfect:

  - Spray all surfaces (hard and soft surfaces) with disinfectant spray. Reapply as needed to achieve required contact time. Allow all surfaces to dry prior to reattaching or reassembling.

  - If any surfaces may contact a user’s mouth rinse with clean water and a new shop towel to remove any chemical residue. Allow to air dry.

  - Follow manufacturer’s instructions for reassembly.

- After cleaning a “batch” of equipment (e.g. loans from 1 user, items from 1 demo) or after handling any heavily soiled items, remove gloves (and any other soiled PPE) and throw in trash.

- Wash your hands with soap and water again. If continuing to clean, put on new gloves.

- Return item to assigned storage area.

## References

- AmTryke 1410 Foot Cycle Early Intervention Assembly Instructions

  - https://ambucs.org/wp-content/uploads/2015/06/1410-Manual-Instructions.pdf